DEVELOPMENT PROGRESS BRIEF

xoExam™ UI/UX Prototype

Tablet controller for the Xenon Ophthalmics™ eye exam system

Currently deployed: v0.1.8  ·  May 18, 2026

Executive summary

v0.1.8 brings the Visual Fields test to full clinical fidelity — the third examination in the suite to reach this standard, after Visual Acuity (v0.1.4) and Color Vision (v0.1.7). This release also folds the test into the unified UI/UX framework that now governs every clinical examination in xoExam: the same exam shell, the same safety-critical cancel behaviour, the same look and feel, the same right-sidebar session controls. The Visual Fields rebuild is the first to be delivered entirely in-place under the unified architecture rather than as a parallel test stream — establishing the pattern for every remaining test rebuild.

The clinical accuracy work in this release reflects the standards of the Humphrey Field Analyzer (HFA), the global gold standard for automated perimetry. The result is a test that a practising ophthalmologist would recognise as clinically complete: industry-standard indices, the Glaucoma Hemifield Test, total deviation and pattern deviation probability plots, age-banded reference values, and proper Hodapp-Anderson-Parrish severity grading.

Snapshot

•   1 release shipped  (v0.1.8)

•   3 tests at full clinical fidelity  (Visual Acuity · Color Vision · Visual Fields)

•   17 application screens  (login, dashboard, exam flows, admin)

•   1 unified examination framework  (ExamShell — same chrome, cancel UX, and session controls across every test in the suite)

Version timeline

v0.1.8   May 18, 2026   CLINICAL ACCURACY + UI/UX INTEGRATION · CURRENT

Visual Fields rebuilt to clinical fidelity and folded into the unified examination framework

Visual Fields became the third test in the system to reach full clinical fidelity. The rebuild covered three connected efforts: (1) clinical accuracy — bringing the test in line with the Humphrey Field Analyzer reference standard; (2) flow correction — the OD → OS eye sequence and binocular-protocol handling that automated perimetry requires; (3) framework integration — folding the test into the same ExamShell, cancel behaviour, and visual system used by every other examination in the suite.

Clinical accuracy

•   Hodapp-Anderson-Parrish severity bands applied throughout, with industry-standard MD/PSD/VFI thresholds driving colour-coded indices and a one-sentence clinical interpretation at each severity level.

•   Glaucoma Hemifield Test (GHT) added as a per-eye categorical pill — the primary categorical output on a real HFA printout. Five categories with colour coding: Within Normal Limits, Borderline, Outside Normal Limits, General Reduction of Sensitivity, Abnormally High Sensitivity.

•   GHT override rule: if any eye returns Outside Normal Limits or General Reduction, the Patient Classification banner is forced up at least one severity band even when the global indices look normal — the standard clinical reading rule for early glaucoma detection.

•   Total Deviation and Pattern Deviation probability plots added to the per-eye detail sections, rendered in the HFA-standard five-tier probability symbology.

•   Age-banded reference MD table added under every report (20–30: −0.3 dB through 80+: −3.6 dB) for direct comparison against age-matched normal populations.

•   Reliability indices rebanded to the HFA’s actual XX-flag thresholds. Standard preset: False Positives <15%, False Negatives <33%, Fixation Losses <20%. A tighter Strict preset (10% / 20% / 15%) is exposed as a practice-level setting.

•   Pattern-specific point counts (10-2: 68 points; 24-2: 54; 30-2: 76; Esterman: 120 etc.) drive the live Responses gauge — previously hard-coded to 54 regardless of pattern.

•   Foveal threshold per eye now displayed in the report under Patient Information and annotated directly on the central dot of each sensitivity diagram.

Eye sequence and flow

•   Clinically correct OD → OS sequence for bilateral monocular protocols. The earlier specification’s OS-first default was clinically backwards — universal medical convention is right-eye-first, and the default has been corrected.

•   Binocular-by-design protocols (Functional Esterman, Modified Esterman, FDP, Combined FDP) correctly run as single OU tests with no eye sequence. There is no monocular equivalent of these patterns clinically — Esterman is a binocular DVLA/DMV driving fitness test by specification.

•   Transition modals between eye phases display patient-positioning instructions ("Cover the patient’s right eye with the occluder…") with a visual breadcrumb of completed, current, and pending eyes.

•   In-test eye picker locks the moment the first response is recorded, preventing accidental mid-test eye reassignment.

UI/UX framework integration

•   The previous six-phase pre-test flow (eye selection → exam settings → pattern → foveal calibration) collapsed under the canonical three-phase ExamShell contract (ready / testing / report) shared by every test in the suite. Pre-test setup steps now live correctly as sub-states under ready, putting the safety-critical cancel behaviour in charge the moment testing begins.

•   Canonical Cancel UX now governs Visual Fields: the back arrow is hidden during active testing to prevent accidental data loss via muscle-memory back clicks; in its place, a labelled red Cancel test button opens the confirmation modal (Continue test default / Cancel test destructive). Same behaviour as Visual Acuity and Color Vision.

•   The bespoke internal cancel modal and custom header chrome have been removed entirely. Visual Fields now matches the rest of the suite in look, feel, and safety behaviour — no longer an outlier.

•   Session notes textarea added to the right sidebar during testing — consistent with Color Vision v0.1.5.

•   Sticky sub-bar at the top of the testing content carries the inline eye picker and breadcrumb — the same pattern established in Color Vision v0.1.7, now adapted to Visual Fields.

Visual and typographic sweep

•   Goldmann spelling corrected throughout. The application previously misspelled "Goldman" in config labels, stimulus-size selectors, and report fields.

•   Sentence case applied to every button label across the test (CONTINUE → Continue, START CALIBRATION → Start calibration, EXPORT EXAM → Export report, SAVE & CLOSE → Certify & close). Brings Visual Fields in line with the rest of the suite’s typographic discipline.

•   Solid accent backgrounds replace gradient buttons, matching the design-system standard set across the rest of the application.

•   Lens stepper buttons raised to a 36×36 pixel minimum touch target. The previous 24-pixel controls were below the tablet’s touch-target requirement.

•   Per-eye report layout follows the cumulative multi-eye structure established by Color Vision v0.1.7: Patient Classification banner with the clinical bottom line, per-eye summary cards in a flexible 1–3 column grid, per-eye detail sections with the sensitivity map plus Total Deviation and Pattern Deviation plots, lens correction used during the test, session duration, and a closing Reference Values + Clinical Interpretation column pair.

New practice-level tweaks exposed

•   Reliability strictness preset (Standard or Strict) — practices can tune to their internal quality-assurance bar.

•   Default start eye (OD or OS) — accommodates left-handed examiners or practice-specific preferences.

Deployment package alignment

•   Three previously-missing files corrected in the deployment manifest: DoctorsPage.jsx (from the v0.1.2 doctor section rebuild), the cropped D-15 launch icon (from v0.1.6), and the xoFit™ and xoLab™ marketing imagery (from v0.1.3) were silently absent from the deploy list. All three are now included.

•   Ishihara plate file references corrected from .png to .jpg to match the actual asset files.

•   Standard downloadable deployment package generated as a single .zip for direct upload to GitHub. 58 files, all version-tagged to v0.1.8.

Open clinical questions flagged for engineering

Four points were flagged in the specification document for resolution with the engineering / firmware team before clinical launch. The application proceeded with appropriate placeholder values so the design pass could complete without blocking.

•   GHT calculation source — the xoExam liquid-lens headset firmware will need to compute and return the categorical GHT result, or the application must compute it cloud-side from raw threshold data. Either path is supported by the current UI.

•   Real liquid-lens correction range — placeholder ranges (SPH ±14.00 D, CYL ±6.00 D) are exposed; the actual headset capability must be confirmed and the steppers clamped accordingly.

•   Point counts for the proprietary Screening and Goldmann / Kinetic patterns — placeholder values used pending engineering confirmation.

•   Foveal calibration output shape — single dB threshold (as currently displayed) or full four-step staircase response array.

Currently deployed

xoExam™ v0.1.8 — Target platform: Android tablet · 1280×800 base, responsive to 1920px.

Tests completed at full clinical fidelity: 3 of 19 — Visual Acuity · Color Vision · Visual Fields.

Application shell: Complete · login → welcome → dashboard → all sections.

Unified examination framework: Complete · every test now mounts inside the same ExamShell with shared chrome, timer, patient header, cancel safety, and finish/certify flow.

What’s next

•   Wavefront Aberrometry — extract from the bundled-tests file and apply the same clinical-accuracy + framework-integration standard established in v0.1.7 and v0.1.8.

•   Ocular Motility — extract from the bundled-tests file and design pass.

•   Test-complete detection banner — design and pilot a subtle banner that signals when a test has reached its natural clinical endpoint, then back-port to all 19 tests.

•   Engineering handoff package — once test rollout completes, prepare the developer handoff for backend integration (Bluetooth / Wi-Fi device events, cloud sync, HIPAA-compliant data storage, authentication, real-time monitoring).

Method Marketing Agency  ·  UI/UX design and prototype          Brief generated May 18, 2026  ·  xoExam™ v0.1.8