# xoExam™ UI/UX — Engineering Handoff Specification

## v0.2.5 · May 22, 2026 · Pupillometry · Beta scope CLOSED

Method Marketing Agency for Xenon Ophthalmics Inc.

## 1. What changed since v0.2.1

This v0.2.5 reissue captures the Pupillometry clinical-fidelity rebuild — the sixth and final beta-scope test. With Pupillometry shipped, the originally-defined beta scope is closed. All sixteen sections of v0.2.1 remain authoritative unless updated below.

Area  ·  v0.2.1  ·  v0.2.5

|---|---|---|

Tests at clinical fidelity  ·  5 of 19  ·  6 of 19 (Pupillometry added)

Beta scope status  ·  One remaining (Pupillometry)  ·  CLOSED

Per-test clinical specs in briefs/  ·  5 (VA, CV, VF, WFA, EOM)  ·  6 (+ Pupillometry)

Component count in components/  ·  25  ·  25 (PupillometryTest.jsx replaced in-place, not added)

Open clinical questions on the kickoff agenda  ·  ~22  ·  ~29 (7 added from Pupillometry §H)

Doctor sign-off pattern  ·  Established (v0.1.9)  ·  Applied to Pupillometry per §6 of Clinical Standards Reference

The post-beta roadmap consists of the remaining 13 visual-fidelity placeholder tests, in priority order to be set with the clinical evaluation team.

## 2. Engineering quick start (unchanged from v0.2.1)

Stack: React 18.3.1 + Babel Standalone 7.29.0 via CDN. No build toolchain. index.html is the single entry point.

Open index.html directly in a modern browser (Chrome / Edge recommended). The prototype renders immediately — no server, no install, no build.

Production target: React Native + Kotlin Android tablet. The .jsx source is structured for mechanical port — inline styles, hooks-only, no exotic dependencies. See the v0.2.1 spec §10 for the full RN portability table.

## 3. ExamShell Component Interface Contract (unchanged)

Every test component must conform to this contract exactly. Pupillometry was the sixth test rebuilt against it.

js

function [TestName]({ onBack, tweaks }) {

const accent = tweaks?.accentColor || '#1f8eff';

// ...

return (

<ExamShell

title="[Test name — sentence case]"

accent={accent}

onBack={onBack}

patientName={patientName}

patientId={patientId}

phase={phase}

elapsed={elapsed}

onBegin={startTest}

onFinish={() => setPhase('report')}

onNewTest={resetTest}

rightPanel={phase === 'testing' ? <RightSidebar /> : null}

>

{phase === 'ready'   && renderReady()}

{phase === 'testing' && renderTesting()}

{phase === 'report'  && renderReport()}

</ExamShell>

);

}

Object.assign(window, { [TestName] });

ExamShell owns the canonical Cancel UX: back arrow hidden during testing, labeled red "Cancel test" button takes over, opens confirm modal. Test components MUST NOT build their own cancel dialogs.

For tests with sub-test state machines inside the testing phase (EOM, Pupillometry), finishAvailable() is the local helper that gates ExamShell's green "Finish & Report" button — only at the last enabled sub-step's completion state.

## 4. Pupillometry — engineering-specific integration touchpoints

The clinical spec (briefs/Pupillometry_Clinical_Spec_v2.md) is the authoritative source. Engineering touchpoints specific to Pupillometry:

### 4.1 Hardware contract — open questions to resolve

These seven items are added to the engineering kickoff agenda (full list in §12). Each is flagged inline in the Clinical Standards Reference §6 and in briefs/Pupillometry_Clinical_Spec_v2.md §H.

#  ·  Question  ·  Affects

|---|---|---|

1  ·  Dynamic data shape from firmware — time-series at ~30 Hz, or summary metrics only?  ·  PUP_LightReflexChart data binding; six summary tiles

2  ·  NPi computation location — firmware-side (NeurOptics convention) or UI/cloud-side?  ·  PUP_computeNPi helper (currently UI-computed for the prototype)

3  ·  Swinging flashlight automation — headset-driven alternation, or technician-driven?  ·  PUP_SwingingScreen behavior (currently UI-driven via Start/Pause stimulus button)

4  ·  Pupil response clicker — pupillometry is passive; confirm clicker is not used  ·  Headset event subscription; clicker disable while in Pupillometry

5  ·  Pharmacologic workflow scope — cocaine/apraclonidine for Horner's, pilocarpine for Adie's  ·  Out of scope for v0.2.5; if planned, follow-on pharmaco--prefixed sub-test

6  ·  Ambient light measurement — real headset-interior lux readback, or assumed calibrated?  ·  PUP_LIGHT_CONDITIONS ambient value source

7  ·  Adaptation time enforcement — firmware-enforced ≥ 5 min dark adaptation, or UI countdown?  ·  Static sub-step phase logic; setup-screen warning copy

### 4.2 Data persistence schema additions

Pupillometry adds these per-eye fields to the per-test results object (see v0.2.1 §6 for the cross-test schema):

js

{

testId: 'pupillometry',

eyesTested: ['OD', 'OS'],   // or ['OD'] / ['OS']

patientAge: 42,

static: {

OD: { scotopic: 6.7, mesopic: 4.9, photopic: 3.0 },

OS: { scotopic: 6.6, mesopic: 4.8, photopic: 3.0 }

},

dynamic: {

OD: {

baselineSize: 5.4, minSize: 3.6, constrictionPct: 33.3,

latencyMs: 240, constrictionMs: 320, redilationMs: 1200,

constrictionVel: 5.6, redilationVel: 1.5,

npi: 3.2                   // computed; firmware-side per Q4.1.2

},

OS: { / same shape / }

},

rapd: {

OD: 'none',                  // 'none' | 'trace' | '1plus' | '2plus' | '3plus' | '4plus'

OS: 'none'

},

notes: 'Free-text session notes',

certifiedBy: '<doctor-id>',     // populated at Certify & close

certifiedAt: '<ISO timestamp>'

}

### 4.3 Patient classification logic — pure-function port targets

The following helpers should port directly to the production codebase, with NPi swapped to firmware-provided when Q4.1.2 is resolved:

PUP_getAgeBand(age) — Winn 1994 age-band lookup

PUP_inRange(value, [lo, hi]) — age-band deviation check

PUP_getAnisocoriaSeverity(diffMm, conditionDelta) — light-vs-dark differential rule

PUP_computeNPi(dynamics) — placeholder formula (swap to firmware passthrough)

PUP_npiSeverity(npi) — NeurOptics threshold bands

PUP_getInterp(eyesTested, results, rapdByEye, ageBand, ranDynamic, ranSwinging) — Patient Classification banner severity

PUP_getClinicalFlag(...) — four-pattern flag matcher (Horner's, CN III, RAPD, NPi)

All pure JS, all live in components/PupillometryTest.jsx.

## 5. Open decisions — engineering kickoff agenda (cumulative, v0.2.5)

The cumulative list from v0.2.1 stands. New items added by Pupillometry are flagged with [P]:

Hardware-to-UI event contract (cross-test) — what events the headset fires, what payload, what the UI subscribes to

Per-test data persistence schema — final shape across all clinical-fidelity tests

Authentication model — login, session, doctor identity

Export mechanism — PDF generation, EHR forwarding, cloud sync

Role-based permissions — Doctor sign-off as a gate, not just a label

Patient identity confirmation step — currently assumed pre-test

Compliance scope — HIPAA, MDR class, FDA 510(k) path

[P] Dynamic data shape from firmware (Q4.1.1)

[P] NPi computation location (Q4.1.2)

[P] Swinging flashlight automation (Q4.1.3)

[P] Pupil response clicker non-use confirmation (Q4.1.4)

[P] Pharmacologic pupillometry workflow scope (Q4.1.5)

[P] Ambient light measurement from headset (Q4.1.6)

[P] Adaptation time enforcement (Q4.1.7)

See v0.2.1 §12 for the full annotation on items 1–7; items 8–14 are detailed in §4.1 above.

## 6. Regulatory acknowledgment (unchanged)

The xoExam UI/UX prototype is not a medical device and has not been cleared by FDA, MDR, or any other regulatory body. The clinical-accuracy work in briefs/ and in the Clinical Standards Reference is a transparency document — a description of what the design adheres to — not a clinical validation report or regulatory submission. Production clinical validation is an engineering and regulatory responsibility outside Method's scope.

The same applies to instrument calibration: the prototype assumes the production headset will deliver calibrated stimulus presentation, accurate color reproduction, valid threshold measurements, and consistent ambient conditions. Whether it does is a hardware question outside the UI/UX layer.

## 7. Per-test specs (current)

#  ·  Test  ·  Spec file in briefs/  ·  Clinical-fidelity version

|---|---|---|---|

1  ·  Visual Acuity  ·  VisualAcuity_Clinical_Spec_v2.md  ·  v0.2.3

2  ·  Color Vision  ·  (in CLAUDE.md + Clinical Standards Reference §2)  ·  v0.1.7

3  ·  Visual Fields  ·  VisualField_Clinical_Spec_v2.md  ·  v0.1.8

4  ·  Wavefront Aberrometry  ·  WavefrontAberrometry_Clinical_Spec_v2.md  ·  v0.1.9

5  ·  Extraocular Motility  ·  ExtraocularMotility_Clinical_Spec_v2.md  ·  v0.2.1

6  ·  Pupillometry  ·  Pupillometry_Clinical_Spec_v2.md  ·  v0.2.5

The remaining 13 tests are visual-fidelity placeholders. Specs will be written as each one is brought to clinical fidelity in post-beta releases.

## 8. Document set issued at v0.2.5

Document  ·  Status

|---|---|

xoExam UI-UX Engineering Handoff Specification v0.2.5.docx  ·  This document — supersedes v0.2.1

xoExam Development Brief 05-22-2026 v0.2.5.docx  ·  New per-release brief

xoExam Development Brief.docx (cumulative)  ·  v0.2.5 entry to be appended

xoExam Clinical Standards Reference.md / .docx  ·  Updated to v0.2.5 — Pupillometry §6 added

xoExam Clinical Evaluation Brief.md / .docx  ·  Updated — Pupillometry §6 added

briefs/Pupillometry_Clinical_Spec_v2.md  ·  Authoritative clinical spec for Pupillometry

CLAUDE.md  ·  v0.2.5 entry appended; beta scope CLOSED

README.md  ·  Status bumped to v0.2.5

_dist_v0.2.5/  ·  Deployment package

## 16. Version history

Spec version  ·  Date  ·  Trigger

|---|---|---|

v0.1.8  ·  May 18, 2026  ·  Visual Fields clinical fidelity + first formal handoff issue

v0.1.9  ·  May 21, 2026  ·  Wavefront Aberrometry clinical fidelity + Doctor sign-off pattern introduced

v0.2.1  ·  May 22, 2026  ·  Extraocular Motility clinical fidelity + naming convention sweep

v0.2.5  ·  May 22, 2026  ·  Pupillometry clinical fidelity — beta scope CLOSED

Older versioned specs (v0.1.8 / v0.1.9 / v0.2.1) are retained in the repository so spec evolution is traceable for MPR / regulatory audit.

Method Marketing Agency · xoExam UI/UX Engineering Handoff Specification · v0.2.5 · May 22, 2026