xoExam™ — Development Brief

v0.2.6 · June 2, 2026 · Wavefront Refraction

Method Marketing Agency for Xenon Ophthalmics Inc.

Executive summary

v0.2.6 adds Wavefront Refraction — the combined digital replacement for the traditional two-instrument refraction workflow (autorefractor + phoropter). It brings the count of tests at clinical fidelity to 7 of 19: Visual Acuity, Color Vision, Visual Fields, Wavefront Aberrometry, Extraocular Motility, Pupillometry, and now Wavefront Refraction.

This is the most clinically consequential test in the suite. Its output — the verified prescription (Rx) — is the gate that releases the job object to the downstream xoFit fitting/finishing workflow. It was prioritized ahead of other v0.2.5-review feedback at CD direction.

Wavefront Refraction is added as a new, parallel test built from a copy of Wavefront Aberrometry, the Refraction test's live eye-scan, and the Visual Acuity chart. It supersedes but does not replace Wavefront Aberrometry — both remain in the catalog per CD direction. The handoff doc set has been reissued at v0.2.6 with a precise per-file changelog.

This program is in active, versioned development. Each release brings one or more tests to clinical fidelity and ships an accurate changelog of exactly what was touched.

What changed

`components/WavefrontRefractionTest.jsx` — new two-stage refraction test (1381 lines, `WFR_`-prefixed)

A new test (id: wavefront-refraction, Refraction category) mapped onto ExamShell's three-phase contract as a four-phase clinical flow entry → objective → subjective → report.

Entry. Two-card selector — Full refraction (objective → subjective, recommended) or Subjective only (skip objective, enter manual starting values). Cards drive navigation; no separate Begin button.

Objective stage (Stage 1 — replaces the autorefractor). Monocular wavefront capture, OD then OS, using the Refraction test's live eye-scan look rather than the aberrometry circular-progress animation — a clinically appropriate dark viewer (the headset interior is dim), iris radial gradient, breathing pupil, accent sweep line, rotating dashed centering ring, and a red alignment dot that turns the backdrop green once locked. Sub-stages init → calibrating → capturing. Readiness collapsibles: Patient alignment, Pupil detection, Patient focus. Each eye documents a scan count (1–3) — the wavefront retries on movement/blink and averages; the count tells the clinician how reliable the measurement is.

Subjective stage (Stage 2 — replaces the phoropter). Liquid-lens digital phoropter refinement, step sequence per eye setup → sphere (MPMVA) → jcc-axis → jcc-power → [mpmva2*] → add. Runs OD fully, then OS. Optotype selector (Letters / Tumbling E / Tumbling C) shared with the standalone Visual Acuity test. Additive VA notation (e.g. 20/25 +4) surfaced in the right panel.

*mpmva2 (a second MPMVA) auto-triggers only if cyl changed ≥ 0.50 D or axis changed ≥ 10° from objective.

Clinical non-negotiables baked in

•  Axis before power (JCC). Enforced by step order — the UI cannot reach power refinement before axis is confirmed.

•  Fogging control. After objective capture, a fog prompt offers +0.75 D with a clinical explanation (relaxes accommodation, prevents over-minusing). Skippable. Starting MPMVA sphere = objective sphere + fog.

•  Sphere compensation note during cylinder power: "For each −0.50 D CYL added, add +0.25 D SPH to maintain the circle of least confusion."

•  Rx format discipline. SPH/CYL signed + 2 dp; AXIS integer + °; ADD signed with D or "—". No exceptions.

•  Monocular objective. No OU averaging. OD → OS via breadcrumb + transition modal.

•  Certify & close is the Rx-certification system event. The Rx is not released downstream until certified — flagged to engineering as a system event, not a UI-only action. Report sub-label: "Doctor sign-off · releases Rx to job object."

•  Scan-count documentation per eye, surfaced in the report.

•  No clinical claims. Deliberate divergence from the other clinical-fidelity tests — NO Patient Classification banner, NO interpretive language. A refraction is a measurement-and-verification workflow, not a screening, so interpretive language is stripped at clinical direction.

Report

•  Navy "Wavefront Refraction Report" title (formal, title case).

•  Patient Information including refraction mode, fog applied, JCC power used.

•  Final Prescription — most prominent (2px accent border): large signed Rx per eye + ADD + BCVA.

•  Two-box lower row: Objective Results (navy measurement table, scans taken) and Subjective Correction (delta from objective → final).

•  Wavefront analysis (objective only): vertical view rail (pupil image / centroid / full wavefront / high-low orders), 3 mm / 5 mm ring overlay toggle, "+" zoom-scope opening a pannable enlarged map, color-scale legend.

•  Clinical Summary — data only.

•  Actions: Export report · Compare · Certify & close (with Doctor sign-off label).

`components/VisualAcuityTest.jsx` — geometric optotypes added

Tumbling E (Snellen E — 3 bars + spine on a 5×5 grid) and Tumbling C (Landolt ring with a 1-unit gap) drawn as proper geometric SVG optotypes via VA_Optotype — one consistent symbol per chart, shrinking down the lines, varying orientation, rather than rotated font glyphs. Tumbling C added as a 4th chart type. The same optotype shape is shared with Wavefront Refraction's subjective stage. This serves patient accessibility (children, non-literate, non-Latin-script readers report orientation rather than naming a glyph).

Retained, not replaced

components/WavefrontAberrometryTest.jsx is unchanged and stays in the catalog. Wavefront Refraction supersedes it as the working refraction tool but does not remove it — per CD direction both tests remain available.

Tests at clinical fidelity (7 of 19)

The entire UI/UX ships as one unified build — v0.2.6. The column below shows *when each test last reached clinical fidelity*, not a per-test version number. A test listed at an earlier release is correct and unchanged since then, not running older software.

| # | Test | Reached clinical fidelity in |
| --- | --- | --- |
| 1 | Visual Acuity | v0.2.3 (Tumbling E/C optotypes added v0.2.6) |
| 2 | Color Vision (Ishihara + D-15 Farnsworth) | v0.1.7 |
| 3 | Visual Fields | v0.1.8 |
| 4 | Wavefront Aberrometry | v0.1.9 |
| 5 | Extraocular Motility | v0.2.1 |
| 6 | Pupillometry | v0.2.5 |
| 7 | Wavefront Refraction | v0.2.6 |

The remaining 12 tests in the catalog are visual-fidelity placeholders awaiting their own clinical rebuilds. The clinical evaluation panel's input on which to prioritize next is welcomed.

Open items for engineering / next review

•  Certify & close → Rx-release system event. MPR to wire this as the trigger that releases the verified Rx to the xoFit job object. Couples to the role/permissions model (only an authorized doctor may certify).

•  VA pass/fail threshold for the subjective stage (WFR_lineIsPassed, currently "more than half correct") — Gary to validate against a published standard before lock.

•  Patient-facing headset/clicker view for objective self-administered mode — out of scope for this build; separate scope to define with MPR.

•  Liquid-lens / eye-tracking capability confirmation for objective readiness panel values and green-lock timing.

Document set issued at v0.2.6

| Document | Status |
| --- | --- |
| xoExam UI-UX Engineering Handoff Specification v0.2.6.docx | New — supersedes v0.2.5 |
| xoExam Development Brief 06-02-2026 v0.2.6.docx | New release-specific brief |
| xoExam Development Brief.docx (cumulative) | Updated with v0.2.6 entry |
| briefs/WavefrontRefraction_Clinical_Spec_v1.md | Authoritative clinical spec for Wavefront Refraction |
| xoExam Clinical Standards Reference v0.2.6.docx / .md | Reissued at v0.2.6 — Wavefront Refraction added |
| xoExam Clinical Evaluation Brief.docx / .md | Reissued at v0.2.6 — Wavefront Refraction added |
| CLAUDE.md | v0.2.6 entry appended; component index + test-stream table updated; naming conventions codified (rules 13–14) |
| README.md | Status bumped to v0.2.6 |

Deployment

•  Code: components/WavefrontRefractionTest.jsx (added); components/VisualAcuityTest.jsx (edited)

•  Title bumped: index.html → v0.2.6

•  Deploy harness: deploy.html → v0.2.6

•  Distribution package: _dist_v0.2.6/

•  Live: [xoexam-uiux.netlify.app](https://xoexam-uiux.netlify.app)

*Method Marketing Agency · xoExam UI/UX Development Brief · v0.2.6 · June 2, 2026*